Evaluating Glycemic Benefits of an AI-driven Glucose Prediction App Via Digital Twin Technology

Background & Aim

AI-driven glucose predictions enable CGM users to manage their glucose levels more proactively. The Accu-Chek® SmartGuide Predict app provides notifications for low glucose levels within 30 minutes (LGP) and nighttime hypoglycemia risk (NLP), as well as a continuous 2-hour glucose forecast (GP). This study evaluates the potential glycemic benefits of using the app through in-silico analysis.

Methods

Using the UVa digital twin methodology, 2458 days from 20 representative individuals from the REPLACE-BG dataset were simulated. Interventions consisted of 15g fast-acting carbs following LGP notifications, 30g slow-absorbing carbs following high NLP risk prediction at 22:30, and advancing/adding correction boli if GP forecast >220 mg/dL. Baseline intervention included a GP-like hyperglycemia mitigation strategy, wherein predictions were substituted with current CGM measurements. All interventions employed hypoglycemia treatments that were aligned with ADA guidelines. Nighttime meals/boli were omitted to better assess NLP.

Results

LGP reduced %time <70mg/dL (4.2±1.5 vs 0.46±0.28 p<1e-4). GP reduced %time >180mg/dL (25.3±10.2 vs 22.0±8.6 p<1e-4). NLP decreased overnight %time <70mg/dL (5.6±2.2 vs 3.8±1.2 p<1e-4). Combining all features significantly reduced time outside range.

Table 1. Mean±Std per 24h/night population outcomes (n=20) for the different evaluated interventions. The night period is defined as 22:30-05:30. Scope; Intervention; Mean CGM (mg/dL); % Time < 70 mg/dL; % Time in [70,180] mg/dL; % Time > 180 mg/dL. 24h; BASELINE; 147.7±16.9; 4.2±1.5; 70.3±9.8; 25.3±10.2. 24h; LGP; 153.8±15.9***; 0.46±0.28***; 72.9±9.9***; 26.5±10.0***. 24h; GP; 143.6±14.5***; 4.4±1.4*; 73.5±8.4***; 22.0±8.6***. Night; BASELINE; 134.5±14.2; 5.6±2.2; 76.2±7.5; 18.1±8.4. Night; NLP; 141.0±12.1***; 3.8±1.2***; 75.4±7.2*; 20.7±7.6***. 24h; LGP+GP+NLP; 150.1±13.4*; 0.51±0.29***; 76.1±8.3***; 23.3±8.4**. * (p < 0.05) ** (p < 0.01) *** (p < 0.0001)

Discussion/Conclusion

The Accu-Chek® SmartGuide Predict app has the potential to enhance glycemic control in T1D. Limitations include the in-silico nature of the study and the assumptions made in the user behaviour models.