Evaluation of CT:VQ Imaging as an Alternative for Lung Perfusion Assessment

Eikelis N1, Hatt CR1, Nilsen K1, Dusting J1, Siddharthan T2, Dwyer A3,4, Sharma D3, Mogel G1, Fouras A1, Ludwig KD1

14DMedical Limited, Melbourne, Australia; 2University of Miami, Miami, Florida, USA; 3SAHMRI; 4National Imaging Facility, Australia

Introduction/Aim:

A novel imaging approach, CT:VQ (4DMedical™), quantifies and generates 3D ventilation (V) and perfusion (Q) images from non-contrast inspiratory/expiratory chest CT scans. The aim of this study is to evaluate the performance of CT:VQ in generating 3D Q images from non-contrast chest CT scans and to compare it directly with single-photon emission CT (SPECT) perfusion images. This comparison was performed in clinical conditions involving patients with subacute pulmonary embolism, chronic lung diseases, and those undergoing lung volume reduction therapy.

Method:

In a prospective study, six experienced radiologists with average of 24.3 years of experience, who routinely review V/Q studies conducted evaluations. Seventy-seven subjects from several clinical sites (including University of Miami and SAHMRI) underwent paired inspiratory/expiratory chest CT and SPECT V/Q scans. Perfusion images were reviewed using an online DICOM viewer for reporting findings within an embedded survey, assessing lung perfusion across six zones on a monotonic scale (1=absent to 5=well perfused). Inter-modality correlation was analyzed using Kendall’s Tau (τ). The similarities in perfusion intensity distributions across modalities were also compared by radiologists in a side-by-side view (1=no agreement to 5=excellent agreement) for each lung zone and reported as mean ± standard deviation (SD).  
Results:

Kendall’s Tau values ranged from 0.702 to 0.765, all statistically significant (p<0.001), indicating strong ordinal agreement between CT:VQ and SPECT across all zones (Table 1). Radiologist-rated perfusion intensity agreement scores were high, with mean ratings above 3 for each reader and exceeding 4 when averaged, signifying good to excellent inter-modality agreement.

Conclusion:

CT:VQ software generates clinically interpretable perfusion images that align closely with SPECT imaging. This supports its role as a viable, non-nuclear alternative for regional lung perfusion assessment.

Table 1:

| Lung Zone | Inter-Modality Reader Agreement (τ) | Radiologist-Averaged Perfusion Rating (Mean ± SD) |
| --- | --- | --- |
| Upper-Left | 0.743 | 4.32 ± 0.95 |
| Middle-Left | 0.758 | 4.32 ± 0.90 |
| Lower-Left | 0.712 | 4.34 ± 0.89 |
| Upper-Right | 0.765 | 4.43 ± 0.86 |
| Middle-Right | 0.734 | 4.38 ± 0.87 |
| Lower-Right | 0.702 | 4.33 ± 0.95 |

Grant Support: This study was supported by a grant from 4DMedical

Declaration of Interest Statement: Eikelis, Hatt, Nilsen, Dusting, Mogel, Fouras and Ludwig are employees of 4DMedical
Key words: SPECT VQ, ventilation perfusion, imaging