PEF Repeatability Criteria for Spirometry – Relative or Absolute?

Glavas, V1, Johnston S1, Johnson K1, Yeo, A1,2, Pendergrast, C3

1Department of Thoracic Medicine, Royal Adelaide Hospital, Adelaide, Australia 2 University of Adelaide, Adelaide, Australia 3 Department of Respiratory & Sleep Medicine, Women’s and Children’s Hospital, Adelaide, Australia

Introduction/Aim: The 2019 joint American Thoracic Society and European Respiratory Society (ATS/ERS) Standardisation of Spirometry guidelines do not specify a recommendation for PEF repeatability. PEF is an important indicator of patient effort. Earlier repeatability recommendations from Quanjer et al. suggest a variation of < 10% in spirometry PEF or 0.67L/s when PEF is used for monitoring. The aim of this study was to determine if adding PEF repeatability criteria (< 0.67 L/s or <10% of maximum) to existing ATS/ERS guidelines reduced patient ability to meet repeatability criteria, increased the number of trials performed in a test session or altered the reported FEV1. Method: We retrospectively reviewed patients >18 years with a PEF of < 6.7L/s who performed acceptable and repeatable pre-bronchodilator spirometry (ATS/ERS, 2019) at the Royal Adelaide Hospital from January 2022 to October 2023. Random stratified sampling identified 100 patients who met the inclusion criteria (n=50 PEF < 3 L/s, n=50 PEF > 3 L/s and < 6.7 L/s). The number of patients meeting repeatability criteria, the number of trials required to reach repeatability criteria, and reported FEV1 were recorded. Results: In patients with a PEF > 3L/s, adding 0.67L/s and 10% PEF repeatability criteria resulted in 98% and 82% achieving 3 repeatable trials respectively. In those with a PEF < 3 L/s, 92%, and 52% respectively, could meet repeatability criteria. In those who could achieve 3 repeatable trials, the number of manoeuvres required did not change between groups. There was no significant difference in reported FEV1 when adding PEF repeatability criteria to the ATS/ERS spirometry guidelines. Conclusion: Patients with low PEF (<3L/s) were more likely to achieve 3 repeatable spirometry trials using absolute (0.67L/s) as opposed to relative (10%) variability criteria with no significant impact on the number of trials performed or the reported FEV1. Key Words: Spirometry, repeatability, PEF, FEV1 Nomination for New Investigator Award No Grant Support: N/A