Minimally Invasive Reduction and Fixation of Sacroiliac Luxations Using Tele-robotic and a Dedicated Novel Instrument System – A Cadaveric Study

INTRODUCTION: SacroIliac Luxation/Fractures (SIL/F) are commonly stabilized using an ilio-sacral lag screw and washer. Open Reduction and Internal Fixation (ORIF), have been associated with postoperative complications, including poor screw alignment and suboptimal fixation.

To address ORIF shortcomings, Minimally Invasive Osteosynthesis (MIO) techniques were developped.1 Standard MIO limitations include reliance on intraoperative imaging (IOI), exposure to ionizing radiations and time-consuming drill guide orientation over the safe sacral corridor to ensure accurate screw fixation.

To reduce radiation exposure, a novel SIL Instrument System (SILIS™) and Minimally Invasive Lucent Aiming Device (MILAD™) was developed. Compared to ORIF, SILIS-MILAD assisted SIL/F treatment, significantly improved accuracy and reliability of sacral screw placement.2,3 Yet this approach may require positional adjustments to maintain spinal orientation and drill guide alignment over the SSC.

Advancements in IOI and robotics have improved implant placement accuracy during elective surgeries such as arthroplasties. While robotic assisted SIL fixation in 3D printed pelvic models showed promise in a feasibility pilot study from our lab (unpublished data), preoperative registration protocols were too cumbersome for clinical application at this time. By virtually eliminating most presurgical processes, tele-robotic could circumvent the shortcomings of our fully automated robotic approach. This assumption however has yet to be tested.

The purpose of this proof-of-concept study is to evaluate the accuracy and reliability of a remotely controlled dual light weight robotic arms (LWRA) used for reduction and fixation of SIL/F in a cadaveric study. We hypothesize that tele-robotic surgery would improve accuracy and repeatability in screw placement compared to previously published data of SILIS-MILAD assisted surgery.

MATERIALS & METHODS: A right SIL/F was created on ten cadavers from dogs that had died for reasons unrelated to this study. Pre-operative radiographs and CT were obtained for surgical planning. Sacroiliac reduction and lag screw fixation were performed under fluoroscopic guidance using two remotely controlled 6-axis LWRA equipped with SILIS-MILAD custom end effectors (Fig-1). Optimal spinal orientation (L7 transverse processes superimposition and L7-S1 end-plate parallelism) was maintained utilizing a remotely controlled 6 DoF Stewart platform fitted with a custom radiolucent table. The surgical team remained at least 4 meters away from the X-ray source throughout SIL/F reduction and MILAD placement. Pilot and glide hole drilling as well as screw insertion were manually performed by the surgeon without fluoroscopy. To reduce metal artifact, aluminium screws were used rather than surgical cortical screws. Based on dog size, screw lengths and diameters (43mmx4mm or 38mmx3.5mm) were chosen to encompass at least 60% of the sacral width as measured on postoperative ventrodorsal radiographs.

Postoperative outcome measures included sacral screw orientation angles in the dorsal (Cranio-Caudal Angle, CCA) and transverse planes (Dorso-Ventral Angle, DVA). As described in previous studies, both CCA and DVA were measured on their respective postoperative CT-based MPR views2,3 (Fig-2). In addition, SIL/F reduction accuracy was evaluated on ventrodorsal radiographs (Fig-3) using the 1) percentage of reduction (PoR) defined one minus the ratio of sacroiliac incongruence length to sacral wing length (100% indicating anatomic reduction)4 and 2) hemipelvic canal width ratio (HCWR) defined as the right hemipelvic canal width (acetabulum medial wall to sagittal plane) divided by the left hemipelvic canal width (one indicating symmetry).5

Angular and anatomic measurements were reported using descriptive statistics (mean±SD). Angular data was compared to historical data from our laboratory.

RESULTS (Table 1): Surgery was performed on 10 cadavers with a mean BW of 14.2±2.8 kg (12.5–33). Mean screw DVA was 0.4°±0.7° (0°–2°). Mean screw CCA was 0.3°±0.7° (0°–2°). Mean PoR was 98.9±0.02% (92.8-100%). Mean HCWR was 0.97±0.1 (0.9-1.0). Mean screw purchase was 84.7%±0.1 (62.5-100%).

DISCUSSION: This is the first study documenting the use of tele-robotic surgery for the reduction and fixation of SIL/F in veterinary orthopedics. Our results demonstrate that mean screw angles in the transverse and dorsal planes were substantially smaller and more consistent than those reported with ORIF (DVA 7.4°±4.9° [1.1°-14.5°] and CCA 8.5°±5.3° [3.3°-21.4°]) and MIO of SIL/F (DVA 1.6°±1.1° [0.3°-3.7°] and CCA 1.2°±0.6° [0.7°-2.1°]). Additionally, reduction was nearly anatomic in all cases as demonstrated by high PoR and HWCR.

One study limitation was radiation exposure was not recorded. However, internal (unpublished) data from SILIS-MILAD assisted surgeries showed undetectable radiation beyond 2 meters from the C-Arm. Another limitation is that DVA and CCA were compared to historical data. Nonetheless, this data was generated from our laboratory while using similar protocols.

Over the past decade, treatment of SIL/F has seen considerable developments from conservative management to ORIF to various MIO approaches. Minimally invasive techniques and novel instrumentation have helped reduce morbidity and improve clinical outcome of this painful and debilitating condition. We believe that tele-robotic surgery is the next step in improving accuracy and reliability of MIO of SIL/F and should eventually be considered as a valid alternative to current approaches.

ACKNOWLEDGEMENTS: While the SILIS-MILAD is patented to the senior author no financial interest are associated with this study.

REFERENCES

1. Tomlinson et al. VS 1999

2. Dejardin et al. VCOT 2016

3. Dejardin et al. VS 2017

4. Froidefond et al. VS 2023

5. Naiman et al. VS 2021