Light Ion-Induced Nuclear Reaction Studies Using the SLICS Detector at the University of Rochester’s Omega Laser Facility

C. J. Forrest,1,b) A. Aprahamian2, G. Hale5, S. Padalino3, M. Paris5, J. P. Knauer1, S. P. Regan1, M. Rosenberg1, W.U. Schroeder1, A. Schwemmlein1, M. Wiescher2, and M. Yuly4

1Laboratory for Laser Energetics, University of Rochester, Rochester, New York, 14627, USA

2University of Notre Dame, South Bend, Indiana, 46556, USA
3SUNY Geneseo, Geneseo, New York, 14454, USA

4Houghton University, Houghton, New York, 14744, USA

5Los Alamos National Laboratory, Los Alamos, New Mexico, 87544, USA

Corresponding Author Email: cforrest@lle.rochester.edu

A new activation detector optimized for short-lived isotopes was designed, installed and tested on OMEGA the Laser Facility. The Short-Lived Isotope Counting System (SLICS) can be positioned as close as 35 cm from the OMEGA target chamber center in a ten-inch-manipulator (TIM). [1] Stable target samples can be mounted on the front end of the detector assembly to undergo triton-, deuteron-, or neutron-induced reactions that produce short-lived reaction products that beta decay. The emitted betas are detected in a phoswich configuration where two layers of plastic scintillators are placed behind the target. The first layer is thin and allows most electrons to reach the subsequent, thick layer. Electrons suffer characteristic energy losses in both layers, allowing distinction from other radiation. Different scintillation decay times make it possible to discriminate the two different signals in the attached photomultiplier tube. Using the existing triton laser-ion acceleration platform on OMEGA, this detector can be used to measure rarely studied tritium-induced reactions on light nuclei. [2] The design of the detector and its performance during a first experiment employing deuterium as a surrogate will be discussed.

[1] A. Schwemmlein et al, Rev. Sci. Instrum. 97, 043301 (2026)

[2] A. Aprahamian et al, J. Phys. G: Nucl. Part. Phys. 52 063001 (2025)

Specify category (Big Bang Nucleosynthesis (BBN), Stellar Nucleosynthesis (SN), Target Normal Sheath Acceleration (TNSA), Tritium-induced nuclear reactions):