Programmable Digital-Based Potentiostat For

Non-Enzymatic Glucose Detection At Microelectrodes

Andrea De Gregorio a, Simone Luigi Marasso b,c,, Salvatore Guastella a,

Andrea Lamberti a,d, Paolo Stefano Crovetti a, Mara Serrapede* a,d

a Politecnico di Torino, Dipartimento di Scienza Applicata e Tecnologia (DISAT), Corso Duca Degli Abruzzi, 24, 10129, Turin, Italy

b IMEM-CNR (Istituto dei Materiali per l'Elettronica ed il Magnetismo), Parco Area delle Scienze, 37/A, 43124 Parma PR, Italy

c PiQuET-Piemonte Quantum Enabling Technology, Str. delle Cacce, 91, 10135 Torino TO, Italy

d Istituto Italiano di Tecnologia, Center for Sustainable Future Technologies, Via Livorno 60, 10144, Turin, Italy

mara.serrapede@polito.it

This work presents an ultra-low voltage and ultralow power chronoamperometric electrochemical sensor readout integrated circuit (IC) in 130nm CMOS for non-enzymatic glucose detection featuring a reconfigurable Digital-Based (DB) Potentiostat. The signal transfer and noise characteristics of the new digital-based architecture are analytically described in the frequency domain for the first time by an equivalent linearized model that is validated by simulations and experiments. Based on experiments, the proposed DB potentiostat enables the detection of a wide electrochemical current range, spanning from 600 pA to 650 nA, with R2=0.991 linearity and consumes only 1.65nW (53.5 nW) at V dd = 300mV ( V dd = 500 mV). The proposed DB readout is tested in a proof-of-concept platform for nonenzymatic glucose detection with nanostructured microelectrodes, demonstrating successful glucose detection at physiological levels at the lowest reported voltage and power, even in the presence of an interferent (ascorbic acid) and under uncontrolled environmental conditions, thus revealing a strong potential for emerging Point of Care (PoC) diagnostics applications.

[1] J. Kim, A. S. Campbell, B. E. F. de A´vila, and J. Wang, “Wearable biosensors for healthcare monitoring,” Nature Biotechnology, vol. 37, pp. 389–406, 2019.

[2] A. Meimandi, G. L. Barbruni, P. S. Crovetti, and S. Carrara, “Sub- 9 × 29 μm2, 1.2nw, fully digital potentiostat design for dopamine sensing,” IEEE Sensors Journal, vol. 25, no. 20, pp. 38792–38799, 2025.

[3] P. S. Crovetti, “A digital-based analog differential circuit,” IEEE Transactions on Circuits and Systems I: Regular Papers, vol. 60, no. 12, pp. 3107–3116, 2013.

[4] S. A. G. Evans, J. M. Elliott, L. M. Andrews, P. N. Bartlett, P. J. Doyle, and G. Denuault, “Detection of hydrogen peroxide at mesoporous platinum microelectrodes,” Analytical Chemistry, vol. 74, no. 6, pp. 1322–1326, 2002.

[5] M. Serrapede, G. Denuault, M. Sosna, G. L. Pesce, and R. J. Ball, “Scanning electrochemical microscopy: Using the potentiometric mode of secm to study the mixed potential arising from two independent redox processes,” Analytical Chemistry, vol. 85, no. 17, pp. 8341–8346, 2013.